Frascati 22-02-2011
TECHNICAL SPECIFICATIONS (provisional) FOR THE IMPLANTATION OF BAKING OF MP

As you can see from the drawings attached the vacuum chamber has a shape almost cylindrical. 

The material of which are the three central areas is aluminium, the two external areas above and below, and the two caps are instead of stainless steel. 

The three central areas have also a layer of insulation already mounted. 

The purpose of the heating system is to perform a gradual degassing water of the vacuum Chamber up to a temperature, uniform as possible, of 90 ° C. 

For this we need the supply and installation of the following materials: 

· heating 

· Insulation Tapes of parts discoveries where possible (without obstruct access to flanges). 

· Provision of temperature sensors 

· provide temperature control system 

· Assembly and installation of the entire system.

Estimation (in the maximum configuration) of materials to be provided for the control of each sector and two caps. 

• 16 tapes circular 1 kW 220Vac 

• 21 temperature sensors, in the Centre of the 5 sectors (5), on the two flanges between an area and the other (10) and on the two caps (6) 

• 16 controlled current regulators (PID) 

This schema can be adjusted to optimize costs and performance, then the unit price must be provided for each component. 
Using your products we ask for an offer about this configuration:

1. 24  heating tapes TeMS6, PN 184 778-000

2. 112 mt of heating cable KTe, PN 480 964-000

3. 21 temperature sensors PT100, PN 068 430-000  

4. 7 temperature controls with RS485 TCONTROLL-CONT-03/COM, PN 1244-00 6982

5. thermal insulation for the two caps and a template cover to sustain it.
